Appendix A:  Vacuum Systems

## Introduction

With proper care, most vacuum systems can operate trouble-free for long periods of time: not only does continuous operation not harm them but it in fact cleanses them of some contaminants.  It should be understood that the ultimate vacuum (within the capability of the system) and the time it takes to reach that vacuum are determined by the cleanliness of the system.  It takes only one “lunkhead” opening a valve at the wrong time to contaminate a vacuum system, slowing the pumping speed and ruining its vacuum capability.  Those operating vacuum systems should read and understand operating instructions before starting the experiment, and ask for assistance if the instructions are not understood.

## Vacuum Systems

Vacuum systems in the Advanced Laboratory consist of a mechanical rotary pump for low vacuum, an air-cooled oil diffusion pump for high vacuum, and the necessary valves and traps to interconnect them.  Low vacuum is measured by thermocouple gauges and high vacuum by cold- cathode ion gauges.

Figure A-1

A mechanical rotary pump is used for both foreline and roughing pumping.  (See Fig. A-1).  A solid steel cylindrical rotor is mounted eccentrically in a larger steel cylinder.  Contact between the two is effected by two spring-loaded vanes in diametrical slots in the rotor.  As the rotor turns, it pulls air particles through the intake port and expels them through the outlet valve.  Vacuum sealing is supplied by the precise machining of the parts and special mechanical pump vacuum oil.

High vacuum is supplied by a three-stage oil diffusion pump.  A special silicone oil is boiled by electric heaters.  As the oil turns into a gas it naturally increases in volume.  The expanding gas jets up the concentric tubes and is deflected downward pulling air particles from above. When the gas hits the cooled walls of the pump, it condenses and returns to the bottom of the pump for reheating, while the air particles are pulled from the diffusion pump through a “foreline” by the mechanical pump.  A diffusion pump cannot operate without this foreline pumping.

Since there are gaseous oil particles in the lower part of the diffusion pump and a high vacuum above, some method must be used to prevent the oil from “backstreaming” into the high vacuum area.  One such method is the use of baffle plates which, being at a lower temperature, condense the oil.  The method we use is a liquid nitrogen cold trap which not only condenses the oil but also freezes out water vapor, air, and other contaminants.

Thermocouple gauges are used to measure low vacuum pressures.  A thermocouple is a junction of two metals which generates a small amount of electric current depending on the temperature of the junction.  In a vacuum thermocouple gauge the junction is heated by a calibrated amount of current fed to a heater adjacent to the junction.  The current generated by the junction is then measured by a sensitive ammeter.  As the vacuum increases and the number of air particles decreases, there is less transfer of heat to the junction, causing less current to be generated, thereby giving an indication of the vacuum.

In cold cathode ion gauges used to measure high vacuums, a high voltage anode ionizes the gases in the gauge, thereby causing a current to flow as an indication of vacuum pressure.  The higher the vacuum, the fewer the gas particles, less ionization, smaller current.  The advantage of a cold cathode gauge is that it is undamaged if accidentally brought up to atmospheric pressure.  The disadvantage is its slowness in reacting to pressure changes.  Cold cathode gauges should never be operated above 25 microns as the ionization of organic compounds at that pressure contaminates the gauge.

Interconnecting the various parts of a vacuum system are special vacuum valves.  See Figure A-2.  Since pumping speed is determined in part by the size of the pumping line, vacuum valves are usually operated either fully opened or fully closed.  The largest valve is the HIGH VACUUM valve connecting the diffusion pump with the experiment.  A “FORELINE” valve connects the lower part of the diffusion pump with the mechanical pump.  A “ROUGHING” valve connects the experiment to the mechanical pump so that the bulk of the air in the experiment can be removed before connection to the diffusion pump.  Small “AIR INLET” valves are installed to permit parts of the equipment to be brought up to atmospheric pressure.  Most removable parts of a vacuum system are vacuum-sealed to the system by “O-Ring” seals, exceptions being the mechanical pump which is connected to rough rubber vacuum tubing, and the bell jar which uses an “L” shaped gasket.  A light coating of vacuum grease is used on these seals to prevent vacuum leaks, but gobs of grease are not necessary or desirable.

Figure A-2

## Operating procedure to achieve high vacuum

### Warming Up The System

- Make sure all valves are closed (turned clockwise).  Little force is needed for these valves to seal, and over-tightening can destroy them.

- Turn on the mechanical pump.  (The switch is located on the control panel.)

- Open the foreline valve.  (Turn the knob counter-clockwise.)

- Turn the knob on the thermocouple gauge controller to TC-1.  (This measures low vacuum in the diffusion pump.)

- When the thermocouple gauge controller reads 35 microns or less, turn on the diffusion pump.  (The switch is located on the control panel.)  It takes approximately 30 minutes for the diffusion pump to reach operating temperature.  If the system has been contaminated, the diffusion pump may need to run for several hours before use.

- Turn on the ion discharge gauge to the lowest scale (25×10-3mm).

### Roughing Out the Bell Jar

- Before proceeding, CHECK:

- That the mechanical pump is running;

- That the foreline valve is the only valve that is open;

- That the bell jar is installed and secured properly.

- Close the foreline valve.  (NEVER leave this valve closed for more than 5 minutes when the diffusion pump is at operating temperature.)

- Open the roughing valve.  (Normally when the mechanical pump is removing air from the system it will make some noise.)

- Turn the thermocouple gauge controller to TC-2.  (This measures low vacuum in the bell jar.)

- When the thermocouple gauge reaches 50 microns or less close the roughing valve.

- Open the foreline valve.  (Remember this valve should not be closed for more than 5 minutes.)

### Pumping with the Complete System to Achieve High Vacuum

- Before proceeding, CHECK:

- That the bell jar has been roughed out (TC-2 reads less than 50 microns);

- That the foreline valve is open;

- That the roughing valve is closed.

- Open the high vacuum valve;

- Pour liquid nitrogen into the cold trap;

- Adjust the scale on the ion gauge controller to read the high vacuum pressure.

## Opening bell jar

- Close the high vacuum valve;

- Turn the thermocouple gauge controller to TC-1;

- Open the air inlet valve;

- When the hissing stops close the air inlet valve;

- Remove the bell jar;

- If you are not finished with the experiment, return to section III.B, “Roughing Out The Bell Jar”.

## After you have finished

Vacuum systems should not just be turned off.  The diffusion pump must be cooled down before the mechanical pump can be turned off.  This is accomplished by a few simple steps:

- Close the high vacuum valve;

- Turn off the ion gauge controller;

- Turn the diffusion pump off;

- Turn off the thermocouple gauge controller;

- The mechanical pump should run for at least 45 minutes as the diffusion pump cools down.  You do not have to wait for the diffusion pump to cool; simply let the mechanical pump run overnight.  This will not hurt it.